## Slide 1: Dynamic Adjusting Prosthetic Socket
Joshua Friscia (RBE/ME), Meagan Hiatt (RBE/ECE), Mollie Myers (BME), Jacob Zizmor (RBE)
Advisors: Gregory Fischer, Cagdas Onal
Co-Advisors: Michael Delph, Christopher Nycz

## Slide 2: Project significance
1.6 million Americans with limb loss (2005)1
Importance of Prostheses
  Return to activities
  Increased mobility
Many challenges with prosthetic use
  Unnatural forces  injury2
  Poor fit  prosthesis misuse2
  Increased energy from user3
Acute erthyema and edema4

## Slide 3: Current Prosthesis
Adjustable
  Mechanical
  Inflatable Inserts
Daily Fit
  Vacuum
  Suction
Static Fit
  Compression Release Stabilized (CRS) Socket
  Conventional Sockets
Socket with Pump [6]
SmartPuck [7]
Notes: There are three main categories for current prosthesis. There are prothesis that are adjustable throughout the day, prosthesis for which the wearer sets the fit when they don the socket, and static fit prostheses.

## Slide 4: Purpose
Controllable smart variable geometry socket
Utilize compression/release stabilized (CRS) socket ideas9
Sense User Activity
Prosthetic Socket7
Adjust in fit8

## Slide 5: Improvements
Bladder system
  Pneumatic vs. hydraulic
Custom electrical system
Advanced sensor systems
Lighter weight/increase portability
Safety Regulations
Previous MQP Project [10]

## Slide 6: Project Objectives
Safe
Must not cause more harm than current socket technologies
Reliable
Must fail in predictable ways which shall not harm user
Comfortable
Be comparable in comfort to current technologies
Meet FDA device requirements11
Minimal user-input
No need for adjustments throughout the day after initial donning
Long-lasting (daily use)
No need to recharge or adjust for minimum of average shopping trip length (58 minutes)12
Compact
Able to be mounted on the socket
Able to be worn under clothing
Lightweight
Within 25% of the current heaviest socket on market

## Slide 7: Research Analysis

## Slide 8: FDA/CDRH Certification
3 Categories
Based upon regulatory control required for safety
Detailed testing data recorded for future certification
Class I
External limb components
Socket components
Class II
Most hearing aids
C-leg microprocessor knee
Class III
Stair climbing wheelchair
Foot drop orthosis
Other limb braces
Powered wheelchair
[11]
[13]
[14]
[15]

## Slide 9: Types of Socket Solutions
Mechanically adjusting
Vacuum Suspension
Inflatable Inserts/Bladders
Compression Release Stabilized (CRS) Socket

## Slide 10: Mechanically Adjusting
Examples:
  LIM16
  Revolimb17
Pros:
  Fully adjustable
Cons:
  Potential to over-tighten and occlude circulation
Notes: Mechanically Adjusted sockets include the LIM and Revolimb. These sockets are fully adjustable through the use of ratcheting straps or tension dials to increase or decrease pressure and fit on the limb. This is beneficial, adding a longer useful lifespan to the limb itself, but requires constant adjustment from the user throughout the day and there is the possibility of overtightening the device, cutting off circulation in the limb.

## Slide 11: Vacuum Suspension
Examples:
  Vacuum Assisted Suspension System (VASS) 18
Pros:
  Better adhesion to socket
Cons:
  Possibility of soft tissue damage if donned incorrectly
Notes: Vacuum suspension is another type of socket solution. It involves the user donning a liner, inserting their limb into the socket, and evacuating the air between the liner and socket. This creates a negative pressure on the limb and actually controls volume loss of the limb, providing better adhesion to the socket. The donning process is, however, much longer and can result in soft tissue damage if proper care is not taken and the socket is donned incorrectly.

## Slide 12: Inflatable Air Inserts
Examples:
  Pump it up! 19
  Pneu-fit20
Pros:
  Easy to use and inflate
Cons:
  Do not work over wide volume range
Notes: Inflatable air bladders can be used as inserts in static sockets to help with volume control throughout the day. They’re pretty easy to use, only requiring that the user pump up the bladder to their desired pressure level. However, air is a compressible fluid and thus the bladders need to be pumped up to higher pressures and they don’t work over a wide volume range

## Slide 13: Inflatable Liquid Bladder
Example:
  Smart Variable Geometry Socket21
Pros:
  Wider volume management range
Cons:
  Requires pumping or hydraulic system
Notes: Inflatable liquid bladders are a much better option for volume control than air bladders. The incompressibility of liquids means that they have a much wider volume management range. They of course need some sort of pumping or hydraulic system to control the amount of fluid in the bladders. One example is the smart variable geometry socket, which utilizes the pumping cycle from walking and a control circuit to add fluid into the bladders. This design, however, does not allow individual control of the bladders and the user must be walking to maintain fluid in the bladders.

## Slide 14: Compression Release Stabilized
Currently in development
Pros:
  Better torque transfer from limb to socket
Cons:
  Static system
Notes: Inflatable liquid bladders are a much better option for volume control than air bladders. The incompressibility of liquids means that they have a much wider volume management range. They of course need some sort of pumping or hydraulic system to control the amount of fluid in the bladders. One example is the smart variable geometry socket, which utilizes the pumping cycle from walking and a control circuit to add fluid into the bladders. This design, however, does not allow individual control of the bladders and the user must be walking to maintain fluid in the bladders.

## Slide 15: Dynamic Adjusting Prosthetic Socket (DAPS)
Our Project Approach
Research
Testing
Design Ideas
Outline of Design
Prototype/ Testing Design
Design Alterations
Final Testing & IRB

## Slide 16: Bladder Types
Variation in shape22-24
  Rectangular
  Cuff
  Cylindrical (tubing)
  Custom
Materials
  Silicon
  Canvas
  Kevlar
  Neoprene
  Polyurethanes
  Nylon
  Etc.

## Slide 17: Electromyograms (EMG)
Research Tool25
Upper limb prosthesis26

## Slide 18: Signal Noise
Electrode inherent noise27
Movement artifact
Electromagnetic noise
Internal noise
Inherent noise

## Slide 19: Noise Removal
Design techniques27,28
    Pre-amplifications
Electrical Filtering29,34
  Why we need it?
  10-20Hz Min.
  400-450Hz Max.

## Slide 20: Noise Removal
Electrical filtering
  Butterworth filter29.30
  Chebyshev filter30
  Elliptical filter30
  Bessel filter30,31
Pass Band
Stop
Band
[30]

## Slide 21: Noise Removal
Electrical filtering
  Butterworth filter29.30
  Chebyshev filter30
  Elliptical filter30
  Bessel filter30,31
  Laguerre filter32,33
    All Pass
[34]
[35]

## Slide 22: Noise Removal
Electrical filtering
  Butterworth filter29.30
  Chebyshev filter30
  Elliptical filter30
  Bessel filter30,31
  Laguerre filter32,33
    All Pass
[34]
[36]

## Slide 23: Noise Removal
Electrical filtering
  Butterworth filter29.30
  Chebyshev filter30
  Elliptical filter30
  Bessel filter30,31
  Laguerre filter32,33
    All Pass
  High/low pass34
  Rectifiers34
[37]
[38]
[39]

## Slide 24: Signal Analysis
Intervals
Amplitude

## Slide 25: Design review
Notes: Based on the sockets from our research, work done on the previous project, and our own project goals, we decided on liquid bladder socket with a hydraulically actuated system.

## Slide 26: Bladder Actuation Design
Why hydraulics?42
  Wider volume range than air
  More precise fluid control
  More independent control of bladder
  Easier to implement than pump-reservoir system
  Can create a “dynamic” CRS control
Notes: Why hydraulics? As stated previously, liquid has a much wider volume management range than air. Hydraulics allow us independently control single bladders or pairs of bladders and provide more precise fluid control. We determined that a hydraulic system would be easier to implement than a pump-reservoir system which would require a series of switches, sensors, and a relatively powerful pump. Lastly, with our proposed hydraulic system, we could create a “dynamic” Compression Release Stabilized socket which can provide better torque transfer as well as managing limb volume and pressure put on the limb throughout the day.

## Slide 27: Bladder Actuation Design
Considerations43
  Actuation of syringes
  Type of fluid
  Waterproofing electronics/leak prevention
  Syringe/bladder volumes
  Sensors
  System weight
Notes: For our hydraulic system, there are a long list of things we need to take into account when building and testing our design. Listed above are just a few including the actuation of the syringes – what force does our motor or linear actuator need to provide and how do we make it non-backdriveable. Type of fluid we’re using is also important, we will be using water for testing as it’s basically free and easy to clean up, but we will also see if there are benefits in using other types of liquids as well. Waterproofing the system and preventing leaks is also important to prevent destroying the electronics and making sure that the system maintains appropriate pressures on the users limb.
Bladder and syringe volumes must be large enough to account for maximum limb volume change throughout the day.
Sensors will be required to sense and control pressure.
And lastly, the end product has to be light enough for an amputee to actually use it, so choosing lightweight components is a very important part of the design as well.

## Slide 28: Hydraulics Diagram
Syringe
Pressure Sensor
Tubing
Bladder
Flex Sensor
Socket Wall
Blow-off valve
Linear actuator/motor
[44]
Notes: Here is a diagram of our proposed hydraulic design for a single bladder. We plan on having between 4-6 total bladders in the system with the same exact setup. A linear actuator or motor drives the syringe piston back and forth, pushing the fluid through the tubing and into the bladder. A force from the limb will be applied to the bladder when it’s placed in the socket and pressure sensor inline with the tubing will detect changes in the system pressure due to the limb force. We also plan on putting a blow off or release valve for safety to make sure the maximum pressure in the system does not reach dangerous levels or cause components to burst. Additionally in the socket itself, we may place flex sensors to detect areas of excess pressure and provide feedback that can be used for analysis by a prosthetist in determining how well the socket has been working.

## Slide 29: Hydraulics Calculations
Force Plate Testing
Motor criteria44
  Meet F1 value: ~40.86N
  Stroke length d1: ~15.4cm
  Adequate speed
Further testing
  Syringe-bladder test rig
FLimb = F2
Syringe
Bladder
Notes: In order to establish some basic requirements for the system, we did a few tests using force plates to determine the maximum possible force a bladder might experience. We took the max force from running, approximately 2000N and estimated values for the bladder surface area, syringe piston length, and piston surface area. Using this and multiplying by a factor of safety of 2, we determined that our motor input force needs to be about 40N. Additionally, the motor must provide a stroke length of about 15cm or 6 inches to travel the full length of the syringe piston and ideally we will use the fasted linear actuator possible. Lastly, we need to do is set up an test-rig to make sure that out design concept is feasible.

## Slide 30: Bladders Prototypes
Blood pressure cuffs
Fabric
Silicone rubber mold
Surgical tubing
[46]
Notes: For hydraulic bladders, we have 4 different prototype designs. The first design is the repurposing of a blood pressure cuff.

## Slide 31: Bladders Prototypes
Blood pressure cuffs
Fabric
Silicone rubber mold
Surgical tubing
[47]
Notes: The second design is based off of fabric; nylon and vinyl.

## Slide 32: Bladders Prototypes
Blood pressure cuffs
Fabric
Silicone rubber mold
Surgical tubing
[48]
Notes: The third is a mold from silicone rubber.

## Slide 33: Bladders Prototypes
Blood pressure cuffs
Fabric
Silicone rubber mold
Surgical tubing
[49]
Notes: The fourth is a bladder from the use of surgical tubing.

## Slide 34: Design Matrix
Quality | Weight
Force (N) | 33%
Cost ($) | 12%
Durability | 12%
Effectiveness (x minutes) | 11%
Manufacturability | 10%
Customization | 10%
Weight (g) | 7%
Volume (mL) | 5%
Notes: Here you can see the design matrix we’ll be using to score each bladder prototype.

## Slide 35: Sensing & Control
Signal Processing
sEMG Sensors
Motor Drivers
Motors
Front Right
Front Left
Back Right
Back Left
Battery
Voltage Regulators
ADC’s
Pressor Sensors
User Controls (Buttons,
LED Display)
[50]

## Slide 36: Filtering Circuit
High and low pass filters
Full wave rectifiers
Low Pass Filter
High Pass Filter
[37-38]

## Slide 37: Filtering Circuit
High and low pass filters
Full wave rectifiers
Full Wave Rectifier
[39]

## Slide 38: System Control
Microcontroller (NXP1764FBD100)
  I/O Pins
  32-bits
  High clock speed
  USART to USB
[51]
Notes: http://www.ebay.com/itm/NXP-ARM-Cortex-M3-LPC1768-Mini-DK2-Development-board-/181045434143?hash=item2a2726171f

## Slide 39: System Control
Motor Control
  Drivers
  Battery requirements
  Voltage regulators
[52]
Notes: http://www.firgelli.com/products.php?id=41

## Slide 40: System Control
User Interface
  Warnings – visual and audible
  Reset
  LCD
    Sleep/wake
    Battery life
  On/off button
  Load/unload
  Debug/calibrate
[53]

## Slide 41: EMG Wiring
Flex wires with coating
Fitting sock with EMGs
[54]

## Slide 42: EMG Wiring
Flex wires with coating
Fitting sock with EMGs
[55]

## Slide 43: EMG Placement
Muscles to control leg extension/flexion
  Abduction and adduction
Superficial muscles
[56]

## Slide 44: Control Algorithms
[57]
Notes: This is the primary control algorithm. In short, the algorithm checks to see if the wearer is taking a step or moving to a new position. It then sets the goal pressure for the bladders and controls their position with PID.

## Slide 45: Control Algorithms
[58]
Notes: This is the reset interrupt. When the reset switch is set, the interrupt will pull the syringes out to their home position. When the switch is turned off, the interrupt exits and the system returns to normal function.

## Slide 46: Safety
Bladders are too pressurized
Sensor information informs system to retract syringes
Notes: The following are three safety scenarios for which we are planning. The first is when the bladders are too pressurized. In this case the sensor information will inform the system to retract the syringes.

## Slide 47: Safety
Bladders are too pressurized
Sensor information informs system to retract syringes
Bladders are too pressurized &
sensors have failed
User triggers the reset switch, syringes retract
Notes: The second scenario is when the bladders are too pressurized and the sensors have failed. In this case the wearer can trigger the reset switch, depressurize the bladders, and remove their limb.

## Slide 48: Safety
Bladders are too pressurized
Sensor information informs system to retract syringes
Bladders are too pressurized &
sensors have failed
User triggers the reset switch, syringes retract
Unforeseen emergency
Blow off valve releases pressure
Notes: The final scenario is an unforeseen emergency. In this case, the blow off valve will release the pressure in the socket and allow the wearer to remove their limb.

## Slide 49: Project Timeline

## Slide 50: A-Term Schedule

## Slide 51: B-Term Schedule

## Slide 52: C- Term Schedule

## Slide 53: Thank you!

## Slide 54: Questions?

## Slide 55: References
2. K. Ziegler-Graham, E. J. MacKenzie, P. L. Ephraim, T. G. Travison, and R. Brookmeyer, “Estimating the Prevalence of Limb Loss in the United States: 2005 to 2050,” Arch. Phys. Med. Rehabil., vol. 89, no. 3, pp. 422–429, Mar. 2008.
A. F. T. Mak, M. Zhang, and D. A. Boone, “State-of-the-art research in lower-limb prosthetic biomechanics- socket interface: A review,” J. Rehabil. Res. Dev., vol. 38, no. 2, pp. 161–174, Apr. 2001.
R. Alley, CP, LP, T. W. Williams, MA, M. J. Albuquerque, CPO, and D. E. Altobelli, MD, “Prosthetic sockets stabilized by alternating areas of tissue compression and release,” J. Rehabil. Dev., vol. 48, no. 6, pp. 679–696, Jun. 2011.
J. T. Highsmith and M. J. Highsmith, “Common skin pathology in LE prosthesis users,” JAAPA Off. J. Am. Acad. Physician Assist., vol. 20, no. 11, pp. 33–36, 47, Nov. 2007.
http://www.amputee-center.com/pump.htm
Smart Pump
R. Alley, CP, LP, T. W. Williams, MA, M. J. Albuquerque, CPO, and D. E. Altobelli, MD, “Prosthetic sockets stabilized by alternating areas of tissue compression W. A. Haines, J. M. Colvin, M. L. Haynes, C. T. Kelley, M. W. Ford, M. W. Groves, and J. A. Denune, “Prosthetic device utilizing electric vacuum pump,” US7947085 B2, 24-May-2011.
W. A. Haines, J. M. Colvin, M. L. Haynes, C. T. Kelley, M. W. Ford, M. W. Groves, and J. A. Denune, “Prosthetic device utilizing electric vacuum pump,” US7947085 B2, 24-May-2011. Previous MQP Project – Smart Socket 2014-15
U.S. Food and Drug Administration Regulation of Prosthetic Research, Development, and Testing
Kornum, N., & Bjerre, M. (n.d.). Grocery E-commerce: Consumer Behaviour and Business Strategies. Retrieved from https://books.google.com/books?id=cwEUj22Dlj8C&pg=PA82&lpg=PA82&dq=average+time+spent+in+grocery+store&source=web&ots=a9Yrom4v3x&sig=B4AOx9KoOWEtJKol9o6r61RS8Fs&hl=en&sa=X&oi=book_result&resnum=10&ct=result#v=onepage&q=average%20time%20spent%20in%20grocery%20store&f=false
COMFORTFLEX SOCKET SYSTEM
C-Leg 4
RHEO-KNEE
http://revolimb.com/
http://liminnovations.com/
http://www.aetna.com/cpb/medical/data/600_699/0630.html
http://www.amputee-center.com/pump.htm
http://www.oandp.com/shop/product.asp?supplier_id=B4AA5F60-C5E1-4DD5-A25F-E81B96DFB789&product_id=7C850D5A-894C-452B-94B0-94A7513D8F2C&department_id=0EA0CF9B-EB15-4377-8B60-ADA00B0FFBDD
http://www.oandp.org/jpo/library/2003_03_107.asp
http://www.bloodpressureatwatch.com/wp-content/uploads/2014/11/Omron-H-003DS-Small-Adult-Blood-Pressure-Cuff-for-IntelliSense-Monitors-0.jpg
http://uobds.com/Locksmith-Tools/
http://2kdk833xvsai3k7qcq368ael.wpengine.netdna-cdn.com/wp-content/uploads/2013/01/Surgical_Tubing_4c7e6cd42f4a3.jpg
EMG’s as a Research tool
EMG in upper limb prostetichs
EMG Noise
Design tech
Butterworth Filter Design. (n.d.). Retrieved from http://www.electronics-tutorials.ws/filter/filter_8.html
Filter Basics: Anti-Aliasing. (n.d.). Retrieved from https://www.maximintegrated.com/en/app-notes/index.mvp/id/928
Wazir, L., Mohammad, B., Muhammad, M., & Syed, S. (2014). Performance Analysis of Analog Butterworth Low Pass Filter as Compared to Chebyshev Type-I Filter, Chebyshev Type-II Filter and Elliptical Filter. Circuits and Systems, 5, 209–216. http://doi.org/http://dx.doi.org/10.4236/cs.2014.59023
Masnadi-Shirazi, M. A., & Aleshams, M. (2003). Laguerre discrete-time filter design. Computers & Electrical Engineers, 29(1), 173–192.
Manosube, A., Koseeyaporn, J., & Wardkein, P. (2014). PLI Cancellation in ECG Signal Based on Adaptive Filter by Using Wiener-Hopf Equation for Providing Initial Condition. Computational and Mathematical Methods in Medicine, 2014, 11. http://doi.org/http://dx.doi.org/10.1155/2014/471409
Transversal Filter drawn by Meagan H.
All-Pass Filter drawn by Meagan H.
Transversal Filter Response, http://www.tc.umn.edu/~erick205/Papers/paper.html
Hihg-Pass Filters, Low-Pass Filters, 2015 http://www.allaboutcircuits.com/textbook/alternating-current/chpt-8/high-pass-filters/
Filter Compare
Transversal Filter
All Pass Filter
Transversal Response
Josh
Josh
Josh
Josh
https://www.flickr.com/photos/107621760@N03/10668630046
Photo Taken by Jake Z. of Vinyl Fabric
Photo taken by Jake Z. of plastic mold
SURGICAL TUBING | Brownells
Outline of Circuit drawn by Meagan H. Photo of NXP LCP176x from http://www.nxp.com/
http://www.ebay.com/itm/NXP-ARM-Cortex-M3-LPC1768-Mini-DK2-Development-board-/181045434143?hash=item2a2726171f
http://www.firgelli.com/products.php?id=41
User Interface outline Drawn by Meagan H.
Flex Wires, http://en.esskabel.de/ffc-fdc/C21/
Prosthetic Sock, http://www.knitrite.com/about.html
Mollie
Main Control Diagram drawn by Jake Z.
Control Interrupt Diagram drawn by Jake Z.

## Slide 56: References
2. K. Ziegler-Graham, E. J. MacKenzie, P. L. Ephraim, T. G. Travison, and R. Brookmeyer, “Estimating the Prevalence of Limb Loss in the United States: 2005 to 2050,” Arch. Phys. Med. Rehabil., vol. 89, no. 3, pp. 422–429, Mar. 2008.
A. F. T. Mak, M. Zhang, and D. A. Boone, “State-of-the-art research in lower-limb prosthetic biomechanics- socket interface: A review,” J. Rehabil. Res. Dev., vol. 38, no. 2, pp. 161–174, Apr. 2001.
R. Alley, CP, LP, T. W. Williams, MA, M. J. Albuquerque, CPO, and D. E. Altobelli, MD, “Prosthetic sockets stabilized by alternating areas of tissue compression and release,” J. Rehabil. Dev., vol. 48, no. 6, pp. 679–696, Jun. 2011.
J. T. Highsmith and M. J. Highsmith, “Common skin pathology in LE prosthesis users,” JAAPA Off. J. Am. Acad. Physician Assist., vol. 20, no. 11, pp. 33–36, 47, Nov. 2007.
http://www.amputee-center.com/pump.htm
Smart Pump
J. T. Highsmith and M. J. Highsmith, “Common skin pathology in LE prosthesis users,” JAAPA Off. J. Am. Acad. Physician Assist., vol. 20, no. 11, pp. 33–36, 47, Nov. 2007.
W. A. Haines, J. M. Colvin, M. L. Haynes, C. T. Kelley, M. W. Ford, M. W. Groves, and J. A. Denune, “Prosthetic device utilizing electric vacuum pump,” US7947085 B2, 24-May-2011.
http://www.oandp.com/articles/2008-02_01.asp
Previous MQP Project – Smart Socket 2014-15
U.S. Food and Drug Administration Regulation of Prosthetic Research, Development, and Testing
Kornum, N., & Bjerre, M. (n.d.). Grocery E-commerce: Consumer Behaviour and Business Strategies. Retrieved from https://books.google.com/books?id=cwEUj22Dlj8C&pg=PA82&lpg=PA82&dq=average+time+spent+in+grocery+store&source=web&ots=a9Yrom4v3x&sig=B4AOx9KoOWEtJKol9o6r61RS8Fs&hl=en&sa=X&oi=book_result&resnum=10&ct=result#v=onepage&q=average%20time%20spent%20in%20grocery%20store&f=false
COMFORTFLEX SOCKET SYSTEM
C-Leg 4
RHEO-KNEE
http://revolimb.com/
http://liminnovations.com/
http://www.aetna.com/cpb/medical/data/600_699/0630.html
http://www.amputee-center.com/pump.htm
http://www.oandp.com/shop/product.asp?supplier_id=B4AA5F60-C5E1-4DD5-A25F-E81B96DFB789&product_id=7C850D5A-894C-452B-94B0-94A7513D8F2C&department_id=0EA0CF9B-EB15-4377-8B60-ADA00B0FFBDD
http://www.oandp.org/jpo/library/2003_03_107.asp
http://www.bloodpressureatwatch.com/wp-content/uploads/2014/11/Omron-H-003DS-Small-Adult-Blood-Pressure-Cuff-for-IntelliSense-Monitors-0.jpg
http://uobds.com/Locksmith-Tools/
http://2kdk833xvsai3k7qcq368ael.wpengine.netdna-cdn.com/wp-content/uploads/2013/01/Surgical_Tubing_4c7e6cd42f4a3.jpg
EMG’s as a Research tool
EMG in upper limb prostetichs
EMG Noise
Design tech
Butterworth Filter Design. (n.d.). Retrieved from http://www.electronics-tutorials.ws/filter/filter_8.html
Filter Basics: Anti-Aliasing. (n.d.). Retrieved from https://www.maximintegrated.com/en/app-notes/index.mvp/id/928
Wazir, L., Mohammad, B., Muhammad, M., & Syed, S. (2014). Performance Analysis of Analog Butterworth Low Pass Filter as Compared to Chebyshev Type-I Filter, Chebyshev Type-II Filter and Elliptical Filter. Circuits and Systems, 5, 209–216. http://doi.org/http://dx.doi.org/10.4236/cs.2014.59023
Masnadi-Shirazi, M. A., & Aleshams, M. (2003). Laguerre discrete-time filter design. Computers & Electrical Engineers, 29(1), 173–192.
Manosube, A., Koseeyaporn, J., & Wardkein, P. (2014). PLI Cancellation in ECG Signal Based on Adaptive Filter by Using Wiener-Hopf Equation for Providing Initial Condition. Computational and Mathematical Methods in Medicine, 2014, 11. http://doi.org/http://dx.doi.org/10.1155/2014/471409
Transversal Filter drawn by Meagan H.
All-Pass Filter drawn by Meagan H.
Transversal Filter Response, http://www.tc.umn.edu/~erick205/Papers/paper.html
Hihg-Pass Filters, Low-Pass Filters, 2015 http://www.allaboutcircuits.com/textbook/alternating-current/chpt-8/high-pass-filters/
Filter Compare
Transversal Filter drawn by Meagan H.
All Pass Filter Filter drawn by Meagan H.
Transversal Response
http://www.oandp.org/jpo/library/2003_03_107.aspConsiderations Outlined by Josh, Jake and Meagan
Hydraulics System Outline drawn by Josh F.
https://www.physicsforums.com/threads/pressure-diameter-and-flow-in-a-pneumatic-system.742188/ Photo Taken by Jake Z. of Vinyl Fabric
Filter drawn by Meagan H.
Photo taken by Jake Z. of plastic mold
SURGICAL TUBING | Brownells
Outline of Circuit drawn by Meagan H. Photo of NXP LCP176x from http://www.nxp.com/
http://www.ebay.com/itm/NXP-ARM-Cortex-M3-LPC1768-Mini-DK2-Development-board-/181045434143?hash=item2a2726171f
http://www.firgelli.com/products.php?id=41
User Interface outline Drawn by Meagan H.
Flex Wires, http://en.esskabel.de/ffc-fdc/C21/
Prosthetic Sock, http://www.knitrite.com/about.html
http://www.rehab.research.va.gov/jour/2013/504/images/pantall504f01lb.jpg
Main Control Diagram drawn by Jake Z.
Control Interrupt Diagram drawn by Jake Z.